## Slide 1
Mass Flow Sensor (4)
Mass Flow Sensor (12)
Oxygen Sensor (6)
Oxygen Sensor analog circuit
AIR
O2
TO MASK
FROM MASK
EXHAUST
Oxygen blender
Bellow
Set LPM
Limit Switch
HEPA FILTER
HMEF FILTER
HEPA FILTER
C02 cell
BUFFER MEMBRANE
02 cell
Set %02
Control PEEP and PIP
Error correction
for lpm
11
1
2
4
5
6
7
8
9
13
14
15
16
17
18
3
12
10
Limit Switch (12)
Touch Sensor
Touch Sensor Driver
TOP
CENTRAL CHANNEL
BOT
LCD and LCD Driver
12V
MOTHERBOARD
Stepper Motor (11)
12V
Buzzer
3,3V
LED
3,3V
Pressure Sensor (7)
3,3V
Pressure Sensor (13)
3,3V
Connector
MCU
3,3V
ELECTRONIC ARCHITECTURE
Pinch Valve Driver (8.1)
12V
Pinch Valve Driver (14.1)
12V
Pinch Valve Driver (11.1)
12V
VSO Driver for %O2 (3.1)
12V
VSO Driver for LPM (2.1)
12V
Solenoid (2)
12V
Solenoid (3)
12V
Solenoid (8)
12V
Solenoid (11)
12V
Solenoid (14)
12V
Analog
High Current Driver
Digital
Sensor
Circuit
Green is a driver PCB
Light Blue is the Cherry PCB
Actuator
Actuator
Red is the Power PCB
Yellow is the MCU PCB
Dark blue is the 02 PCB
Gray is the motherboard
White is a loose component
Physical housing section
Power Electronics
MOTHERBOARD
Stepper Motor Driver (11.1)
12-24V
Battery Management
PSU
24V
Battery Pack
14,8V
Mezzanine onto Mother
Cherry Blossom
5V
Cherry Blossom
HDMI Stacker
Cherry Blossom
Wi-Fi Stacker
19-08-2020

## Slide 2
Mass Flow Sensor (4)
Mass Flow Sensor (12)
Oxygen Sensor (6)
Oxygen Sensor analog circuit
AIR
O2
TO MASK
FROM MASK
EXHAUST
Oxygen blender
Bellow
Set LPM
Limit Switch
HEPA FILTER
HMEF FILTER
HEPA FILTER
C02 cell
BUFFER MEMBRANE
02 cell
Set %02
Control PEEP and PIP
Error correction
for lpm
11
1
2
4
5
6
7
8
9
13
14
15
16
17
18
3
12
10
Limit Switch (12)
Touch Sensor
Touch Sensor Driver
TOP
CENTRAL CHANNEL
BOT
LCD and LCD Driver
12V
MOTHERBOARD
Stepper Motor (11)
12V
Stepper Motor Driver (11.1)
12-24V
Battery Management
PSU
24V
Battery Pack
14,8V
Buzzer
3,3V
LED
3,3V
LOOMS AND WIRING
Pinch Valve Driver (8.1)
12V
Pinch Valve Driver (14.1)
12V
Pinch Valve Driver (11.1)
12V
VSO Driver for %O2 (3.1)
12V
VSO Driver for LPM (2.1)
12V
Solenoid (2)
12V
Solenoid (3)
12V
Solenoid (8)
12V
Solenoid (11)
12V
Solenoid (14)
12V
Analog
High Current Driver
Digital
Sensor
Circuit
Actuator
UART
Gray is the motherboard
White is a loose component
Physical housing section
Power
MOTHERBOARD
I2C
19-08-2020
Mezzanine onto Mother
Cherry Blossom
5V
Cherry Blossom
HDMI Stacker
Cherry Blossom
Wi-Fi Stacker
Connector
MCU
3,3V
Pressure Sensor (7)
3,3V
Pressure Sensor (13)
3,3V

## Slide 3
SYSTEM SUPPLY
10,5 – 24VDC
BATTERY CHARGE
CONTROLLER
STEPPER MOTOR DRIVER
12VDC-48VDC
ISOLATED AC/DC
220VAC – 24VDC
150W
LEAD CRYSTAL BATTERY
10,5-13,5VDC
STEPPER MOTOR
FLOATING
L
N
E
IDEAL
DIODE
IDEAL
DIODE
POWER SUPPLY ARCHITECTURE
19-08-2020
Filter
Short circuit protection – Fuses only?

## Slide 4
Limit Switch (12)
Touch Sensor
Touch Sensor Driver
TOP
CENTRAL CHANNEL
BOT
LCD and LCD Driver
12V
MOTHERBOARD
Stepper Motor (11)
12V
Buzzer
3,3V
LED
3,3V
Connector
MCU
3,3V
POWER DISTRIBUTION AND GROUNDING
MOTHERBOARD
Stepper Motor Driver (11.1)
12-24V
220VAC to 24VDC
Mezzanine onto Mother
Cherry Blossom
5V
Cherry Blossom
HDMI Stacker
Cherry Blossom
Wi-Fi Stacker
19-08-2020
Solenoid Driver Board
12V
Connector Stackers
Switching DC-DC
24VDC to 12VDC
Linear DC-DC
12VDC to 3,3VDC
L
N
E
Filter
Linear DC-DC
5VDC to 3,3VDC
Switching DC-DC
24VDC to 12VDC
Linear DC-DC
12VDC to 5VDC
HYBRID GROUNDING: SINGLE POINT AND MULTI POINT
Use of linear regulators
Low EMI and cheaper.
Low power consumption on boards makes the lower inefficiency acceptable.
Common-ground impedance can result in noise.
Gauge of ground wire?
Is this a daisy chain ground?

## Slide 5
Touch Sensor
Touch Sensor Driver
TOP
CENTRAL CHANNEL
BOT
LCD and LCD Driver
12V
MOTHERBOARD
Buzzer
3,3V
LED
3,3V
Connector
MCU
3,3V
POWER DISTRIBUTION AND GROUNDING
MOTHERBOARD
220VAC to 24VDC
Mezzanine onto Mother
Cherry Blossom
5V
Cherry Blossom
HDMI Stacker
Cherry Blossom
Wi-Fi Stacker
19-08-2020
Solenoid Driver Board
12V
Connector Stackers
Isolated Switching DC-DC
24VDC to 12VDC
Linear DC-DC
12VDC to 3,3VDC
L
N
E
Filter
Linear DC-DC
5VDC to 3,3VDC
Isolated Switching DC-DC
24VDC to 12VDC
Linear DC-DC
12VDC to 5VDC
HYBRID GROUNDING: SINGLE POINT AND MULTI POINT
Use of linear regulators
Low EMI and cheaper.
Low power consumption on boards makes the lower inefficiency acceptable.
Common-ground impedance can result in noise.
Gauge of ground wire?
Should this be an isolated DC/DC to prevent ground loop?
Should this be an isolated DC/DC to prevent ground loop?
Filter
Stepper Motor Driver (11.1)
12-24V
Stepper Motor (11)
12V